## Slide 1: Fine tuning of the positron beam line
Andrei covali
ENGINEER at GDED, ELI-NP
Coordinator:
CSII Dr.Habil. Nikolay Djourelov
Young Researchers and Young Engineers Days, February 2025

## Slide 2
Summary
Laboratory’s placement and layout
Positron beam purpose and motivation
Tuning and characterization of the beam
Results
Conclusion/Future work

## Slide 3
Laboratory’s placement and layout
Positron Laboratory

## Slide 4
Positron beam purpose
/ pair annihilate (usually)
Two 180°-angle gamma rays are emitted
Gamma rays provide information (among other things) on material’s defects
Annihilation process
Interaction between positron beam with a material
Thermalization = loss of energy
energy < energy
Variance in the annihilation radiation = info on the
All this leads to broadening of the 511keV gamma quanta
Thus, Coincidence Doppler Broadening Spectroscopy was developed

## Slide 5
Gate valves – used to maintain optimal vacuum level
MCP(micro channel plate) with phosphor screen – used to generate a light spot representing the beam spot
CCD camera – used together with a 45°-angle mirror to see the MCP’s output image
Motivation for fine tuning process of the beam line
Fixed MCP is used for Positron Annihilation Lifetime Spectroscopy (PALS)
Start signal: Fixed MCP
Stop signal: Scintillator
How the current through the coils affects the beam

## Slide 6
Problem arises
How can we know precisely where the beam will be
No MCP or other means to center the beam on the sample, except using detectors
Motivation for fine tuning process of the beam line
Deviation to the entry point is precise: 2.5mm/A
Beam spot at the end is a black box (complex behavior after ExB filter)
What is the correlation between CorrCoill Current [A] and the deviation of the beam at the sample [mm]
ExB filter – used for deviating the beam back to the middle
Accelerator – used because our beam is moderated (slow positrons at 29eV)

## Slide 7
Combination of parameters for centering:
Voltage on accelerator’s electrodes (V_acc)
Beam spot size: 3,5,8 mm (adjustable)
Vertical displacement (Vert_displ)
Horizontal displacement (Hor_displ)
For beam centering we used 2 HPGe detectors and 1 Scintillator
V_acc = 1 kV
Vert_displ = 0 mm
Vert_displ = 5 mm
V_acc = 24 kV
Vert_displ = 0 mm
Vert_displ = 5 mm
Fine tuning process of the beam line

## Slide 8
COMSOL Simulations
Vertical displacement behavior
Horizontal displacement behavior
Using COMSOL we predict the behavior of the beam with respect to shifting the center at the beginning on X,Y and X+Y axis, respectively.
Vertical + Horizontal displacement
The beam does not move intuitively:
At the beginning, moving on X => at the end, diagonal shift on primary diagonal
At the beginning, moving on Y => at the end, diagonal shit on the secondary diagonal
At the beginning, moving on both => at the end, straight beam shifting

## Slide 9
Behavior of the beam line (cps) with respect to the current in the correction coils (A)
Difficulty of the task is amplified by the fact that the beam is not a perfect circle
Results of the measurements

## Slide 10
Results of the measurements
We know the number of counts from the detector
We know what % of the beam is passing throw the hole
Direct correlation to the beam spot position
Beam sample: = 10 mm
Beam spot: 3,5,8 mm (adjustable)
For example, let’s analyze and compare 5 mm and 8 mm aperture beam:
Direct relation between the number of counts and the current through the correction coils
Plateau for approximate current of 1 A
Results:
2.5 mm = 0.25 A
The result varies from the previous correlation between beam displacement and coil’s current (2.5 mm = 1 A)
5 mm case:
5mm
8mm

## Slide 11
We found a direct relation between the correction coil’s current and beam spot at the sample!
Combining COMSOL simulations and the experimental results we can confidently conclude that we can manipulate the beam with great precision through the correction coils’ current.
Conclusion
/V | 3 mm | 5 mm | 8 mm
1 kV | 6.11 / 1.5 A | 6.07 / 1.51 A | 6.16 / 1.28 A
10 kV | 6.12 / 1.5 A | 6.02 / 1.47 A | 6.19 / 1.41 A
20 kV | 6 / 1.4 A | 6.01 / 1.47 A | 6.25 / 1.47 A
 | 3 mm | 5 mm | 8 mm
1 kV | 6.11 / 1.5 A | 6.07 / 1.51 A | 6.16 / 1.28 A
10 kV | 6.12 / 1.5 A | 6.02 / 1.47 A | 6.19 / 1.41 A
20 kV | 6 / 1.4 A | 6.01 / 1.47 A | 6.25 / 1.47 A
Future work
Values for the required current though the coils
Upgrading the LabVIEW software for the DAQ to change the proper parameters when the aperture (3,5,8 mm) in changed

## Slide 12
Young Researchers and Young Engineers Days, February 2025
Thank you for your attention!
Questions?
Notes: 12